| Inspection protocol |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Product: | EX-TEC® PM 580/550/500 | Serial no.: |  |  |  |
| Test gas: | ExTox IR | e. g.: | 050 | 30 | 0501 |
| Gas type: | Methane |  |  |  |  |

|  | Date |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | day |  |  |  |  |  |  |  |
|  | month |  |  |  |  |  |  |  |
|  | year |  |  |  |  |  |  |  |

| 1 | Device condition |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1.1 | Perfect condition |  |  |  |  |  |  |  |
| 1.2 | Fine dust filter OK1 |  |  |  |  |  |  |  |
| 1.3 | Disposable/rechargeable battery charge |  |  |  |  |  |  |  |

| 2 | Pump check |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2.1 | Pump error F100 in seal1 |  |  |  |  |  |  |  |

| 3 | Structure |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3.1 | Zero point Test gas: clean air |  |  |  |  |  |  |  |
| 3.2 | Sensitivity Test gas: 1000 ppm CH4 (Target: 1000 ppm, permissible: 700 ppm – 0.14 % vol.) |  |  |  |  |  |  |  |

| 4 | Measuring |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.1 | Zero point CH4 Test gas: clean air (Target: 0 % vol., permissible: -1.0 – 1.0 % vol.) |  |  |  |  |  |  |  |
| 4.2 | Sensitivity Test gas: 100 % vol. CH4 (Target: 100 % vol., permissible: 95 – 105 % vol.) |  |  |  |  |  |  |  |

| 5 | Warning |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 5.1 | Zero point ¦ Sensitivity O2 Test gas: clean air |  |  |  |  |  |  |  |
|  | CH4 (Target: 0 % LEL/% vol., permissible: -3 – 3 % LEL or -0.15 – 0.15 % vol.) |  |  |  |  |  |  |  |
|  | CO (Target: 0 ppm, permissible: -3 – 3 ppm) |  |  |  |  |  |  |  |
|  | H2S (Target: 0 ppm, permissible: -3 – 3 ppm) |  |  |  |  |  |  |  |
|  | CO2 (Target: 0.04 % vol., permissible: -0.02 – 0.10 % vol.) |  |  |  |  |  |  |  |
|  | O2 (Target: 20.9 % vol., permissible: 19.9 – 21.9 % vol.) |  |  |  |  |  |  |  |
| 5.2 | Sensitivity ¦ Zero point O2 Test gas: ExTox IR |  |  |  |  |  |  |  |
|  | CH4 (Target: 2.20 % vol., permissible: 45 - 55 % LEL or 2.00 – 2.40 % vol.) |  |  |  |  |  |  |  |
|  | CO (Target: 40 ppm, permissible: 36 – 44 ppm) |  |  |  |  |  |  |  |
|  | H2S (Target: 40 ppm, permissible: 36 – 44 ppm) |  |  |  |  |  |  |  |
|  | CO2 (Target: 2.00 % vol., permissible: 1.80 – 2.20 % vol.) |  |  |  |  |  |  |  |
|  | O2 (Target: 0.0 % vol., permissible: -1.0 – 1.0 % vol.) |  |  |  |  |  |  |  |
| 5.3 | Visual alarm1 |  |  |  |  |  |  |  |
| 5.4 | Audible alarm1 |  |  |  |  |  |  |  |

| 6 | Comments |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | For example: Housing damaged Filter changed Servicing necessary … |  |  |  |  |  |  |  |

|  | Signature |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

For example: Y for ‘Yes’, N for ‘No’

For example: ¼

Enter current measurement values (actual values) at zero point and sensitivity.